The IMACS Center Field Guide Camera

There are three guider cameras mounted on the Forward Optics Support (FOS) Disk. The Center Field Guide Camera (CFGC) and two off-axis guiders, the Principal (PGC) and Shack-Hartmann (SHGC) Guide Cameras. The cameras have a 105” x 105” square acquisition field. The Center Field Guider is equipped with a slit mask for on-axis single object spectroscopy.

The Center Field Guide Camera, the primary multifunction field acquisition camera, is mounted and fixed to a machined surface on the FOS disk. The camera has three positions to which it can be commanded; out, center (in), and slit view (single object) . The camera is comprised of two main subassemblies; the motion stage and the optics assembly. There are two optical modes available, direct imaging and Shack-Hartmann mode.

Description of Mechanisms

The Motion Stage: The CFGC moves along a linear rail and a linear rail lead screw assembly driven with a five-phase stepper motor. There are two proximity sensors mounted on the main drive rail which provide 'soft limit' end-of-travel and 'home' position feedback. Two microswitches, mounted on the assembly baseplate, provide 'hard limit' end-of-travel position feedback. In this assembly the out position end-of-travel 'soft limit' is also the 'home' position. The main support structure of the camera assembly is mounted to four linear rail blocks, over specified to ensure adequate stiffness, positioning, and runout tolerances.

The Optics Assembly: The main components of the optics assembly include the optics support assembly, optics barrel assembly and filter assembly. The optical paths provide direct imaging/slit view and Shack-Hartmann imaging modes. The barrel subassemblies are mounted to a high precision linear motion stage which is mounted to the optics support assembly. The filter assembly houses two filters, red and neutral density, and is also mounted to the optics support assembly between the main optics barrel assembly and the guider camera. Air cylinders are employed to switch the optics and hold them in place against fixed locating screws. Proximity sensors mounted on the air cylinders are used for feedback. The camera, filter and barrels are located against or positioned with parallel pins.

Two barrels house the optics required for each viewing mode. The direct imaging mode barrel consists of a Canon EF50 lens with manually adjustable focus, achromatic field lens and fold mirror. The pick-off fold mirror is actually a slitmask of highly polished stainless steel with 5 slits of various widths (0.25”, 0.5”, 0.75”, 1”, 1.5”) which can be used for single object slit viewing. A hole feature is also cut into the mask. The Shack-Hartmann barrel houses a Hartmann lenslette array, collimating lens, pinhole baffle and pickoff mirror.

Important Design Considerations

Optics Alignment: Whenever possible the alignment of optics relies on machining tolerances and dead reckoning. When higher stiffness and locating accuracy was needed parallel pins were used where required.

Position Feedback: Because of the high stiffness, repeatability and accuracy inherent in the lead screw drive system, a linear encoder is not used in this system. As such, there is no unique 'home' position, but instead the reverse end-of-travel 'soft limit' and 'home' position are the same.

End-Of-Travel Limits: Proximity sensors are used to mark 'soft' end-of-travel limits. Soft limits are signals to the software to indicate that the end-of-travel has been reached. Microswitches are provided to mark 'hard' end-of-travel limits. In the event that motion continues through the soft limits, the hard limits trigger firmware logic which cause the stepper-motor controllers to stop the cut power to the drive motor.

Shack-Hartmann Pick-Off Mirror: To ensure that only light from a guide star is acquired, the Shack-Hartmann fold mirror is actually just a mirrored spot, roughly 2mm in diameter, on a glass plate.

Shack-Hartmann Template: The Center Field Guide Camera is not equipped with an on board LED for acquiring a Shack-Hartmann template. A special LED mount was designed for this purpose.

Initialization Procedure

Proper initialization and operation of the CFG is possible when the following lockout conditions are satisfied:

The limit switches and sensors must be reporting proper state conditions (they are selected and wired for fail-safe operation).
The optics are moved to the 'home' configuration. 'SHG' optics are in the active position.

With these conditions met, the initialization procedure will 'home' the guide camera by moving it in the reverse direction until the reverse soft limit proximity sensor is tripped. This is the 'out' position as reported in the IMACS GUI.

Normal/Expected Operation

Upon successful initialization the CFGC will be ready for use. The camera should move between limits quickly (under 20 seconds). A low frequency noise can be heard at the beginning and ending for a move. This noise is a characteristic of the harmonics of the device and motor setting. The motor break is normally on when not moving and power must be supplied to the motor to release the brake. The air pressure for all guide cameras should be set to 70 psi. The guide camera regulators are mounted together in a bank near the CFG (above the assembly when IMACS is in the home position).

Before moving into the field-of-view the path is checked and cleared if necessary. The checks and moves are automated and messages may be displayed to inform the user of the moves. When commanded to be moved into the field-of-view any slitmask (or IFU) is removed and the center field slit mask is inserted (if needed). Also, the Principal Guide Camera has a range of motion that overlaps with the CFGC space and it will be moved to allow the CFGC a clear path. In some cases the instrument will be returned to the configuration it was in before the CFGC inserted.

The CFGC will automatically be moved out to the home position if a mask of IFU insert is commanded when the CFGC is in the field-of-view.

Maintenance Considerations and Issues

Routine maintenance should include:

- checking the air pressure – air solenoid requires 35 psi minimum to operate correctly
- check random screws periodically – ensure screws are tightly torqued
- attention to 'normal' sounds – sound can be a good diagnostic of normal operation

Troubleshooting

Several issues have arisen during the assembly and testing and commissioning phases that are addressed for future reference.

Initialization Failure

The Guide Cameras may fail to initialize due to a 'home not found' error. The mostly likely cause for this is that a hard limit switch has been tripped. Resetting the hard limit logic should rectify the problem.

Air Pressure: The pneumatic solenoids that change the state of the optics assemblies requires a minimum of 35 psi for proper operation.

Loss of Power: Power loss to any sensors (due to problems at the rack, severed cables, etc.) will flag software lockouts and prevent actions that could damage the system. Also, the pneumatic solenoids would not function unless powered.

Diagnosing Problems and Troubleshooting Error Messages

The recommendation and procedures provided below are not exclusive or exhaustive.

If an error message occurs during operation the recommended diagnostic procedure is as follows:

1. Note and record the last command attempted.
2. Note and record the state of the instrument as reported by the user GUIs and telescope operator.

The next steps require a break in observing but are necessary for proper and full diagnostics.

1. Check the simple things first: has the rack power been lost; has the rack power cycled (if so the rack computer has been reset – as shown on the rack monitor); has the hard-stop logic been tripped; has the instrument air pressure dropped.
2. Remove the panels and investigate the error of failure. Piece the error message and system state together. Look for problem conditions, what happened and why.
3. Contact key IMACS personnel at SBS; Alan Dressler, Tyson Hare, Christoph Birk, or Alan Baggish.

The following is a list of common error message and what may cause them to be flagged.

home not found: This message would be reported after an initialization command when the wheel isn’t moving. The most probable causes for this problem include: no rack power; the hard limit logic has been tripped; the encoder isn’t operational.