IMACS Guide Cameras

The IMACS Guide Cameras (GCs) were designed to provide on instrument guiding functionality to IMACS. This functionality was needed to due the large field of view exploited by IMACS.

There are three guider cameras mounted on the Forward Optics Support (FOS) Disk. The Center Field Guide Camera (CFGC) and two off-axis guiders, known as the Principal (PGC) and Shack-Hartmann (SHGC) Guide Cameras. The cameras have a 105” x 105” square acquisition field. The Center Field Guider is equipped with a slit mask for on-axis single object spectroscopy. Both off-axis guiders can sweep through a 40 degree arc at a field radius of 15 arcmin, while the Shack-Hartmann has an addition radial motion allowing for a guide star acquisition field extending from 14.5 to 16 arcmin through the entire angular range.

Description of Mechanism

The guide cameras have two main subassemblies; the main drive system and the optical subassembly. Both guide cameras utilize the same main drive system and many common parts, and are similar in function but are sufficiently different in mechanical setup that a separate description of each will given after the common systems are covered.

Main Drive System:

The Principal Guide Camera

The Optical Assembly

The Off-Axis Shack-Hartmann Guide Camera

The Radial Drive System

The Optical Assembly

Important Design Considerations

Normal/Expected Operation

Maintenance Issues

Troubleshooting

Diagnosing and Troubleshooting Error Messages

The recommendation and procedures provided below are not exclusive or exhaustive.

If an error message occurs during operation the recommended diagnostic procedure is as follows:

1. Note and record the last command attempted.
2. Note and record the state of the instrument as reported by the user GUIs and telescope operator.

The next steps require a break in observing but are necessary for proper and full diagnostics.

1. Check the simple things first: has the rack power been lost; has the rack power cycled (if so the rack computer has been reset – as shown on the rack monitor); has the hard-stop logic been tripped; has the instrument air pressure dropped.
2. Remove the panels to investigate the error or failure. Piece the error message and system state together. Look for problem conditions, what happened and why.
3. Contact key IMACS personnel at SBS; Alan Dressler, Tyson Hare, Christoph Birk, or Alan Bagish.

The following is a list error message and what may cause them to be flagged.

Home not found: This message would be reported after an initialization command because the grating isn’t moving. The most probable causes for this problem include: no rack power; the hard limit logic has been tripped; the encoder isn’t operational.